# Protocol for Expansion of PSC Aggregates in the PBS-Mini Vert Bioreactor Driven By Vertical-Wheel® Technology

## Objective

This protocol provides step-by-step guidance for forming and expanding pluripotent stem cell (PSC) aggregates using 2D adherent culture in T-flasks as a seed train, followed by 3D suspension culture expansion in PBS-Mini Vertical-Wheel® bioreactors. The recommended culture parameters serve as a starting point for optimization with various PSC lines and media formulations.

## Materials

### For 2D Culture: Seed Train in T-Flask

Cryopreserved PSCs

T-flasks

PSC culture medium (e.g., mTeSR1, TeSR AOF, StemFlex, StemScale, E8, with mTeSR1 and StemFlex used in optimized processes)

ROCK inhibitor (Y-27632 or equivalent)

Coating substrates (e.g., Matrigel, iMatrix-511, Laminin 521)

Serological pipettes

Aspirating pipettes

15 mL and 50 mL conical tubes

Microcentrifuge tubes

### For 3D Culture: Expansion in PBS-Mini

PSC culture medium (e.g., mTeSR1, TeSR AOF, StemFlex, StemScale; mTeSR1 and StemFlex used in optimized processes)

ROCK inhibitor (Y-27632 or equivalent)

PBS-Mini 0.1 and/or 0.5 vessels

Serological pipettes

Aspirating pipettes

15 mL and 50 mL conical tubes

Microcentrifuge tubes

Enzymatic dissociation reagents (e.g., Accutase, TrypLE)

6-well or 12-well flat-bottom plates

P1000 micropipette tips

Orbital shaker (for conical tube dissociation)

## Procedure

### 1. Overview

Perform expansion starting with 2D adherent culture in T-flasks as a seed train, followed by 3D suspension culture in PBS-Mini Vertical-Wheel bioreactors.

Use recommended culture parameters as a starting point, optimizing as needed based on specific PSC lines and media formulations.

Consult PBS Biotech’s Application Engineers for process tailoring by contacting app.eng@pbsbiotech.com.

---

### 2D Culture: Seed Train in T-Flask

#### A. Inoculation in T-Flask

Formulate PSC Thawing Medium by adding 10 µM ROCK inhibitor to standard PSC culture medium.

Follow vendor recommendations for medium expiration.

Aliquot medium while cold to avoid repeated heating/cooling cycles.

Prepare T-flasks with coating substrate as per vendor instructions.

Add appropriate volume of PSC Thawing Medium to each flask and place flasks in the incubator (37°C, 5% CO₂, 85-95% humidity).

Thaw cryovial containing frozen PSCs in a 37°C water bath until only a small ice crystal remains.

Transfer thawed cell suspension gently into a 15 mL conical tube containing 9 mL pre-warmed PSC Thawing Medium for a total of 10 mL.

Gently triturate the cell suspension 2 times at a 45° tube angle to mix evenly.

Centrifuge cells at 300 x g for 5 minutes.

Aspirate supernatant carefully without disturbing cell pellet.

Resuspend pellet in 5 mL PSC Thawing Medium by gentle trituration 5 times until cells are fully resuspended.

Add another 5 mL PSC Thawing Medium to reach 10 mL total volume.

Triturate suspension 3-5 times; take two 0.5 mL samples in separate 1.5 mL tubes for counting.

Keep cell suspension in incubator while counting.

Vortex samples for 3 seconds at 3000 rpm and perform two counts per aliquot; calculate average of four counts.

Calculate seeding density based on counts; if variance >10% between samples, repeat vortex and counting.

Prepare inoculum at recommended seeding density of 15,000 viable cells/cm² (e.g., for T-25 flask: 187,500 viable cells in 8 mL total volume; for T-75 flask: 562,500 viable cells in 24 mL total volume).

Mix inoculum at least 5 times with serological pipette before seeding.

Add inoculum slowly beneath flask liquid level and gently rock flask in figure-8 motion for even cell distribution.

Confirm even dispersal with microscopy and place flasks into incubator.

#### B. Medium Exchange in T-Flask

Prewarm formulated PSC culture medium.

Inspect cell morphology and confluency under microscope; take images at 4x and 10x magnification.

Transfer T-flasks to biosafety cabinet (BSC).

Aspirate spent medium carefully.

Tilt flask and dispense fresh, prewarmed medium along flask wall to avoid disturbing cells.

Return flasks to incubator.

Perform daily 100% medium exchange with medium volume of 0.32 mL/cm² surface area.

#### C. Harvest from T-Flask

Formulate PSC Inoculation Culture Medium by adding 10 µM ROCK inhibitor to PSC culture medium.

Aliquot enzymatic dissociation reagent (e.g., Accutase or TrypLE), add 10 µM ROCK inhibitor, and prewarm to 37°C.

Confirm cell morphology and confluency (70–80%) under microscope; take images at 4x and 10x.

Transfer T-flasks to BSC; aspirate spent medium.

Add prewarmed DPBS without calcium and magnesium at 0.08 mL/cm² surface area by tilting flask and dispensing along wall.

Gently rock flasks to rinse; aspirate DPBS.

Repeat DPBS wash once more (two total rinses).

Add prewarmed enzymatic dissociation reagent at 0.08 mL/cm² surface area.

Rock flasks gently to cover cells completely.

Incubate flasks at 37°C for 5 minutes; confirm >95% cell detachment microscopically.

If detachment incomplete, incubate additional 1–2 minutes.

If issues persist, wash 1–2 times with DPBS prior to enzyme or extend incubation time.

Collect detached cells with serological pipette into 50 mL conical tubes.

Add PSC Inoculation Culture Medium at 0.08 mL/cm² to rinse and collect remaining cells; add rinse to conical tubes.

Centrifuge cells at 500 x g for 5 minutes.

Aspirate supernatant carefully without disturbing pellet.

Resuspend cell pellet in 10 mL prewarmed PSC Inoculation Culture Medium by gentle trituration.

Triturate 5 times; take two 0.5 mL samples in separate 1.5 mL tubes for counting.

Keep suspension in incubator while counting.

Vortex samples for 3 seconds at 3000 rpm; perform two counts per aliquot. Calculate average from four counts.

Refer to "Inoculation in T-Flask" section for subsequent seeding or to "Inoculation in PBS-Mini" for 3D culture.

---

### 3D Culture: Expansion in PBS-Mini

#### General Parameters for 7-Day PSC Expansion (Table 2)

PBS-Mini 0.1 vessel:

Agitation rate: 60 RPM (scalable range 50 to 80 RPM)

Seeding density: 20,000 cells/mL

Working volume: 60 to 100 mL

Medium exchange: Discrete 50% medium exchange by settling on Days 3, 4, 5, and 6

PBS-Mini 0.5 vessel:

Agitation rate: 40 RPM (scalable range 24 to 42 RPM)

Seeding density: 20,000 cells/mL

Working volume: 300 to 500 mL

Medium exchange: Continuous perfusion (0.5 to 1 vessel volume per day)

---

#### A. Inoculation in PBS-Mini

Decontaminate and install PBS-Mini Base(s) in the incubator; do not store Base inside incubator when idle to prolong lifespan.

Formulate PSC Inoculation Culture Medium by adding 10 µM ROCK inhibitor to PSC culture medium.

Prewarm inoculation medium and batch culture vessels with appropriate volume (e.g., 100 mL for PBS-Mini 0.1, 500 mL for PBS-Mini 0.5) on bases inside incubator.

Equilibrate temperature and dissolved gases overnight with agitation, or minimum 2 hours for 0.1 vessel / 4 hours for 0.5 vessel.

Formulate enzymatic dissociation reagent with 10 µM ROCK inhibitor and prewarm.

Prepare inoculum suspension at target seeding density (20,000 cells/mL).

PBS-Mini 0.1: 400,000 viable cells in total 100 mL volume (95 mL medium + 5 mL inoculum).

PBS-Mini 0.5: 10,000,000 viable cells in total 500 mL volume (475 mL medium + 25 mL inoculum).

Mix inoculum suspension at least 5 times with serological pipette.

Add inoculum suspension to prewarmed vessels.

Place vessels on bases inside incubator; turn on agitation and verify impeller rotation.

After 2 to 4 hours, optionally sample to check cell density and viability.

If viability <80%, review cell bank quality or inoculum handling.

---

#### B. Medium Exchange in PBS-Mini

Formulate PSC culture medium (without ROCK inhibitor for exchanges after inoculation).

Prewarm aliquot of PSC culture medium overnight.

Perform medium exchange on Days 3, 4, 5, and 6 for 7-day culture, exchanging 50% volume.

Turn off vessel agitation; remove vessel and place on incubator shelf.

Allow cells to settle for 5 minutes; tap vessel gently after 2 minutes to dislodge cells from impeller.

Transfer vessel to BSC carefully to avoid resuspending settled aggregates.

Slowly aspirate 50% of spent medium from the top without disturbing aggregates.

To reduce cell loss, centrifuge spent medium and enumerate cells.

Add fresh prewarmed medium equal to volume removed.

Return vessel to base in the incubator and resume agitation.

Repeat exchange for each vessel individually.

---

#### C. Sampling From PBS-Mini

Prepare aliquots of enzymatic dissociation reagent (1.5 mL per sample) and PSC culture medium (1–3 mL per sample); prewarm for 40–60 minutes.

Place a PBS-Mini Base inside BSC for sampling.

Transfer vessel from incubator to BSC; mount on base.

Agitate at 110–120% of culture speed (e.g., if culture is 60 rpm, sample at 66–72 rpm).

Mix culture for 10–20 seconds to ensure homogeneity.

Insert 5 mL serological pipette through vessel port to tip near impeller.

Withdraw samples as follows, avoiding multiple samples from one pipette draw:

Morphology: 1 mL to 6- or 12-well flat-bottom plate.

pH measurement: suitable volume to tube.

Metabolite analysis: suitable volume to tube.

Dissociation and cell counting: 3 mL to 15 mL conical tube.

For morphology, swirl plate gently to center aggregates and image at 4x and 10x within 5 minutes.

Take ≥3 images across cluster.

For pH, measure immediately with calibrated meter.

For metabolite, analyze soon or store at –20°C up to 1 month or –80°C up to 1 year; avoid refreezing thawed samples.

For dissociation, centrifuge 15 mL conical tube at 500 x g for 5 minutes.

Remove supernatant carefully; add 1 mL enzymatic dissociation reagent to pellet.

Resuspend gently to break aggregates without bubbles.

Incubate in 37°C water bath for at least 20 minutes.

Triturate with 1 mL serological pipette until fully dissociated.

Quench enzyme by adding equal volume of PSC culture medium.

Triturate 3 times with P1000 pipette; take two 500 µL samples for counting.

Vortex samples 3 seconds at 3000 rpm; perform two counts each and calculate average from four counts.

Ensure cell viability ≥85%; if lower, review handling and culture parameters.

---

#### D. Harvest and Dissociation in Conical Tube

Formulate PSC Passaging Medium by adding 10 µM ROCK inhibitor to PSC culture medium.

Aliquot enzymatic dissociation reagent with 10 µM ROCK inhibitor; prewarm.

Transfer a PBS-Mini vessel onto BSC base set to 110–120% culture agitation speed (e.g., 66–72 rpm).

Use a 25 mL serological pipette to collect aggregates and medium by inserting vertically near impeller.

Collect entire culture volume into conical tube; gently pipette 1–2 times if needed to resuspend aggregates.

For last 5–10 mL, use 5 mL pipette to access space between impeller and vessel wall.

Centrifuge at 500 x g for 5 minutes.

Aspirate supernatant carefully.

Add enzymatic dissociation reagent as per Table 3:

PBS-Mini 0.1: 20 mL Accutase; up to 150 mL Passaging Medium.

PBS-Mini 0.5: 100 mL Accutase; up to 750 mL Passaging Medium.

Resuspend pellet gently; transfer suspension to flat-bottom bottle.

Place bottle on orbital shaker at 70 rpm in incubator for 20 minutes.

Triturate gently with 10 mL pipette until aggregates are no longer visible; check microscopically.

If incomplete, incubate additional 5 minutes.

Transfer cell suspension to conical tubes.

Rinse bottle with equal volume of PSC Passaging Medium; transfer rinse to tubes to quench enzyme.

Centrifuge tubes at 500 x g for 5 minutes.

Aspirate supernatant carefully; resuspend pellet in PSC Passaging Medium or cryopreservation medium at 2x10^5 to 2x10^6 cells/mL.

Triturate 3–5 times; take two 1 mL samples for counting.

Vortex samples 3 seconds at 3000 rpm; perform two counts per sample and calculate averages.

Confirm viability ≥85%.

For cryopreservation, resuspend cells at desired density (1x10^6 to 1x10^7 cells/mL).

For continued expansion, refer to inoculation in PBS-Mini.

---

#### E. Harvest and Dissociation in PBS-Mini Vessel (In-Vessel Dissociation)

Formulate PSC Passaging Medium by adding 10 µM ROCK inhibitor.

Aliquot enzymatic dissociation reagent with 10 µM ROCK inhibitor; prewarm.

Use volumes as per Table 4:

PBS-Mini 0.1: 20 mL Accutase; up to 150 mL Passaging Medium; agitation speed 80 RPM for dissociation.

PBS-Mini 0.5: 100 mL Accutase; up to 750 mL Passaging Medium; agitation speed 40 RPM for dissociation.

Turn off agitation; remove vessel and place on incubator shelf.

Allow cells to settle for 5 minutes; tap vessel gently after 2 minutes.

Transfer vessel to BSC carefully.

Remove spent medium with serological pipette(s) (10 mL or 5 mL) without disturbing aggregates.

If desired, add DPBS wash before dissociation reagent.

Add prewarmed enzymatic dissociation reagent at volumes per Table 4.

Place vessel on base in incubator; start agitation at recommended dissociation speed.

Note: impeller will be exposed to air, increasing shear to assist dissociation.

After 8 minutes, transfer vessel to BSC and triturate gently with serological pipette until no visible aggregates remain.

Confirm dissociation microscopically; if incomplete, incubate additional 5 minutes.

Transfer suspension to conical tube.

Rinse vessel with equal volume of PSC Passaging Medium; add rinse to tube to quench enzyme.

Centrifuge at 500 x g for 5 minutes.

Aspirate supernatant carefully.

Resuspend pellet with PSC Passaging Medium or cryopreservation medium at 2x10^5 to 2x10^6 cells/mL.

Triturate 3–5 times; take two 1 mL samples.

Vortex 3 seconds at 3000 rpm; count twice per sample and average four counts.

Confirm viability ≥85%.

For cryopreservation, resuspend cells at desired density (1x10^6 to 1x10^7 cells/mL).

For continued expansion, refer to inoculation in PBS-Mini.

---

## Review Checklist

[ ] Growth conditions verified (temperature 37°C, 5% CO₂, 85-95% humidity, incubation times)

[ ] Media composition and volumes verified (PSC culture medium with or without 10 µM ROCK inhibitor)

[ ] Mixing parameters verified (agitation rates: 60 RPM for PBS-Mini 0.1; 40 RPM for PBS-Mini 0.5)

[ ] Incubation times verified (e.g., enzymatic dissociation times, settling times)

[ ] Centrifugation settings verified (300 x g and 500 x g for 5 minutes, per step)

[ ] Step order confirmed against source protocol

[ ] Any ambiguous values flagged with [CHECK]

## Review Flags

## Critical Parameters Checklist (from the protocol)

Temperature: 37°C

CO₂: 5%

Humidity: 85–95%

ROCK inhibitor concentration: 10 µM

2D Culture Seeding density: 15,000 viable cells/cm²

2D Passage 1 Culture duration: 3 days

2D Passage 2+ Culture duration: 4 days

2D Medium volume to surface area: 0.32 mL/cm²

DPBS and enzyme volume to surface area: 0.08 mL/cm²

Enzyme incubation time (2D): 5 min (extend 1–2 min if needed, max 20 min)

Centrifugation settings: 300 x g for 5 min (thaw), 500 x g for 5 min (others)

3D PBS-Mini 0.1 vessel:

Agitation rate: 60 RPM (range 50–80 RPM)

Seeding density: 20,000 cells/mL

Working volume: 60–100 mL

Medium exchange: discrete 50% on days 3–6

3D PBS-Mini 0.5 vessel:

Agitation rate: 40 RPM (range 24–42 RPM)

Seeding density: 20,000 cells/mL

Working volume: 300–500 mL

Medium exchange: continuous perfusion (0.5 to 1 VVD)

Medium equilibration times: overnight preferred; minimum 2 h (0.1 vessel), 4 h (0.5 vessel)

Aggregate settling time during medium exchange: 5 minutes

Sampling agitation speed: 110–120% of culture RPM

Enzymatic dissociation reagent volumes (Accutase):

PBS-Mini 0.1: 20 mL enzyme + up to 150 mL passaging medium

PBS-Mini 0.5: 100 mL enzyme + up to 750 mL passaging medium

Orbital shaker speed (conical tube dissociation): 70 RPM for 20 min

In-vessel enzymatic dissociation agitation speed:

PBS-Mini 0.1: 80 RPM

PBS-Mini 0.5: 40 RPM

Enzymatic dissociation incubation times:

Conical tube: 20 min +/- 5 min extension

In-vessel: 8 min + optional 5 min extension

Cell viability target: ≥85%

Sample vortex conditions: 3 sec at 3000 rpm

Sample triturations (counting): multiple sets of 2 counts per sample, averaging 4 counts total

## Potential Missing Parameters (compare to source)

Thaw step centrifugation speed/time: protocol states 300 x g for 5 min (source states this specifically only for thaw centrifugation), protocol includes it correctly. No missing here.

Enzyme incubation time variability depending on media formulation or coating substrate is detailed in the source but only briefly addressed in protocol; risks of incomplete detachment in 2D are noted only generally.

Source mentions "Do not re-use enzymatic dissociation reagent warmed to 37°C" explicitly in multiple places; protocol references "Aliquot and prewarm" but omits caution about re-use/disposal. [CHECK]

Source includes warnings about not dispensing medium or DPBS directly onto cells but along flask walls; protocol notes this only for inoculation but may omit details for DPBS washes during 2D harvest.

Medium volume per surface area for 2D daily culture medium exchange (0.32 mL/cm²) is indicated in Critical Parameters but the protocol does not explicitly restate the volume in the medium exchange steps. [CHECK]

Source explicitly mentions minimum 2 passages recommended in 2D culture before transitioning to suspension but this is absent in generated text. [CHECK]

Source emphasizes use of aliquoted medium stored cold to avoid repeated heating; in protocol mentioned once but could be underspecified.

Source mentions taking morphology images at 4x and 10x magnification during culture and before harvest; protocol includes imaging steps but does not refer clearly to magnification values or guidelines on abnormalities to look for. [CHECK]

Source advises no medium exchange on D1 or D2 in 3D culture for low inoculation densities; this is absent in protocol. [CHECK]

Source states not to use ROCKi in medium exchanges post-inoculation; protocol includes this but with less emphasis/justification.

## Possible Unsupported Claims

None detected. The generated protocol closely follows the source text and does not appear to contain claims unsupported by the source.

## Step Order / Omission Risks

Protocol omits explicit mention of the recommended minimum of two 2D passages before transition to 3D culture, which is important for downstream performance.

The source includes optional but recommended step of sampling cell viability after 2 to 4 hours post-inoculation in PBS-Mini vessels to verify density and viability; protocol includes this but phrases it as optional/strongly recommended, consistent with source.

In 3D medium exchange, the source recommends doing one vessel at a time to reduce risk of cross-contamination and minimize aggregate settling time; the protocol preserves this but should stress it explicitly as a contamination risk management measure.

Source instructs not to dispense multiple samples from one serological pipette draw when sampling PBS-Mini vessels; protocol preserves this.

Source repeatedly cautions avoiding bubbles during trituration; protocol also includes this.

Source recommends gentle rocking/flask handling during 2D seeding and medium addition to avoid cell disturbance; while protocol captures rocking during inoculation, it could better highlight careful addition of medium during medium exchange steps for cell protection. [CHECK]

Source mentions the risk of aggregate merging if morphology imaging delayed >5 minutes post-sampling; the protocol states images should be taken within 5 minutes, so retained correctly.

Source mentions measuring pH of spent medium to monitor cell density and health; protocol includes this but does not provide detailed procedure for pH measurement at each step as per source. [CHECK]

Possible omission of instructions for washing enzymatic dissociation reagents pre-use or recommended not to reuse warmed dissociation reagents consistently in protocol could affect viability outcomes.

Overall, no major step order reversals observed.

---

# Summary

The generated protocol faithfully incorporates most key numerical parameters and procedural steps from the source text with generally correct order and detail.

However, conservative flags for potential risks from missing or underspecified points include:

Absence of clear recommendation for two 2D passages before 3D transition.

Limited or missing caution against reuse of warmed enzymatic reagents.

Less emphasis or omission of some volume scaling notes (e.g., 0.32 mL/cm² for medium volume in 2D daily exchange).

Limited explicit details about avoiding direct addition of medium/DPBS onto cells in washing/exchange steps.

Missing cautions about not exchanging media before day 3 in 3D if low inoculation density.

Underspecified pH measurement guidance for medium monitoring.

Limited explicit warnings for contamination risk management during medium exchange (single vessels at a time).

Reduced detail on morphology imaging magnifications and abnormalities observation in 2D culture phases.

These flagged omissions and less emphasized warnings could pose risks to reproducibility or cell culture quality if not supplemented with operator experience or further training.