# BioBLU® 0.3c/f Single-Use Bioreactor Instructions for Use

## Objective

To provide detailed user instructions for configuration, operation, and safe handling of the BioBLU 0.3c/f single-use bioreactor and its integration with the DASbox parallel mini bioreactor system.

## Materials

BioBLU 0.3c/f single-use bioreactor (cell culture version: 0.3c/sc; microbiology version: 0.3f)

Reusable accessories (not included in delivery package):

Overhead drive (Type DASbox SU)

Exhaust condenser with adapter plate

DO sensor (DASGIP D 4.7 mm, L 162 mm)

Temperature sensor (Pt100 RTD)

Standard pH sensor (glass, 120/12 mm; Pg 13.5)

Sterile culture medium (100 mL for cell culture; 65-70 mL for microbiology)

Sterile gloves

Ethanol (for disinfection)

Luer lock syringes for sampling

Luer connectors and tubing for media and gas supply

Cleanbench or sterile workbench

DASbox Autoclavable Carrier (for transport)

DASware Control Software

Allen key (1.5 mm) for adjusting magnetic coupling

## Procedure

### 1. Configuration

1.1 Inspect the single-use bioreactor, which is delivered fully assembled with integrated ports for sensors, gassing, media addition, and liquid removal.

1.2 Identify the ports on the head plate labeled as follows:

DO sensor port

Thermowell for Pt100 RTD temperature sensor

Sampling port

Mounting for exhaust condenser

Exhaust gas line

Liquid addition ports (LA(s) submerged, LA(o) overlay)

Pg 13.5 ports for pH or additional sensors

Gas ports (Gas(s) submerged, Gas(o) overlay)

Magnetic coupling for overhead drive

1.3 Note internal fixtures:

Cell culture version (BioBLU 0.3c): 3-blade pitched blade impeller at 3x 45°

Microbiology version (BioBLU 0.3f): two 6-bladed Rushton type impellers

Dip tubes for sampling, harvesting, gassing, and liquid addition

### 2. Preparation and Cultivation

2.1 Typical workflow (perform steps under cleanbench or sterile conditions):

**Cell Culture:**

Unpack bioreactor (4.2.2).

Fill with sterile culture medium via Pg 13.5 port (approx. 100 mL) (4.2.3).

Install calibrated, pre-sterilized pH sensor via Pg 13.5 port (4.2.4).

Add inoculum by pipette or pump.

Connect reservoir bottles with liquids (pH adjustment, feed) (4.2.9).

Transport bioreactors and reservoirs to DASbox with Autoclavable Carrier.

Integrate DO sensor and temperature sensor (4.2.6).

Connect pH sensor if not already connected.

Mount exhaust condenser (4.2.7).

Connect gas supply (4.2.9).

Mount overhead drive (4.2.8).

Adjust cultivation settings and perform DO calibration (refer DASware Control User Manual).

Start cultivation process (refer DASware Control User Manual).

Dispose of bioreactor after cultivation (4.2.11).

**Microbiology:**

Unpack bioreactor (4.2.2).

Fill with sterile culture medium via Pg 13.5 port (65-70 mL) (4.2.3).

Install calibrated, pre-sterilized pH sensor (4.2.4).

Connect reservoir bottles with liquids (4.2.9).

Transport to DASbox (4.2.1).

Integrate DO sensor and temperature sensor (4.2.6).

Connect pH sensor (4.2.4).

Mount exhaust condenser (4.2.7).

Connect gas supply (4.2.9).

Mount overhead drive (4.2.8).

Adjust cultivation settings and perform DO calibration (refer DASware Control User Manual).

Inoculate via sampling valve or under sterile bench, repeat steps 5–10 if inoculation done under bench.

Start cultivation process (refer DASware Control User Manual).

Dispose of bioreactor after cultivation (4.2.11).

### 2.2 Unpacking

Check packaging for damage.

Wear gloves.

Spray packaging with ethanol before placing under cleanbench.

Open sterile double packing by tearing groove at top under cleanbench to avoid contamination.

### 2.3 Filling Culture Medium

Unscrew and remove white blind plug from Pg 13.5 port.

Fill sterilized culture medium via pipette, pump, or funnel: approx. 100 mL (cell culture) or 65–70 mL (microbiology).

Close Pg 13.5 port with blind plug or insert calibrated and sterilized pH sensor.

### 2.4 Installing pH Sensors

Calibrate pH sensor per DASware Control User Manual.

Sterilize sensor (e.g., autoclaving).

Screw sensor into Pg 13.5 port under cleanbench.

### 2.5 Installing Optical Fiber for Optical pH Spot

Use sensor length compatible with bioreactor version.

Insert optical fiber at "pH" labeled cap until it snaps into place, ensuring contact with inner bottom of pH cap.

Avoid excessive force to prevent displacing pH cap.

Connect optical fiber to controller as per controller manual.

Note sensor measurement range: official pH 5.5 – 8.0, recommended 6.5 – 7.5.

Avoid exposing optical pH spot to pH outside sensor range or to direct sunlight/UV, low osmolarity media, solvents, and fluorescent substances.

### 2.6 Integrating DO and Temperature Sensors

Store DO sensor in humidified atmosphere upright in black screw box with sponge wetted with deionized water.

Before insertion, add a few drops of water into DO sensor port to wet the membrane.

Insert DO sensor fully until handle engages, stretching permeable membrane 2–3 mm.

Insert Pt100 RTD temperature sensor carefully into thermowell to the stop without force to avoid damage.

Never activate temperature control without RTD sensor inserted and the bioreactor filled to minimum working volume.

Handle sensors carefully to avoid membrane or cap damage.

### 2.7 Assembling Exhaust Condenser

Insert condenser adapter into holder on head plate gently to avoid cracking.

Press preinstalled exhaust gas tube into groove of condenser adapter.

Push tube insulation onto condenser adapter to stop.

Slide actuator onto magnetic plate; correct position activates cooling automatically.

Remove actuator to deactivate cooling after cultivation.

Position bioreactor in DASbox well before assembling off-gas condenser to prevent tipping due to top-heaviness.

### 2.8 Assembling Overhead Drive

Do not place overhead drive on magnetic coupling while drive is operating to avoid damage.

Position bioreactor in DASbox well for stability due to increased top-heaviness.

Place overhead drive on magnetic coupling in center of head plate, aligning slots for proper position.

### 2.9 Connecting Dispensing Media and Gas Supply

Remove luer caps and connect tubing aseptically to liquid addition ports.

Insert DASbox PA tubes (AD 4 mm) into luer ports of inlet gas filters.

Ensure secure, straight fitting of connections to prevent leakages.

Do not use corrosive, explosive, flammable, or toxic gases.

Do not exceed maximum pressure of 0.4 bar; never block outlet gas tubes or filters.

Confirm exhaust filter ventilation and use overpressure valves if applicable.

### 2.10 Sampling

Remove screw cap and spray sampling valve with ethanol.

Attach luer lock syringe; compression opens valve.

Withdraw sample carefully.

Disinfect valve with ethanol and close with luer cap after sampling.

Do not penetrate valve with needles.

Remove first 1 mL dead volume when performing repeat sampling.

Sampling valve also usable for inoculation.

### 2.11 Proper Disposal

Disassemble reusable accessories.

Empty bioreactor and place in autoclave bag.

Do NOT autoclave bioreactor with culture broth or filled.

Follow national regulations for disposal of biologically active substances.

Wear personal protective equipment (PPE) and protective gloves when handling.

### 3. Troubleshooting

3.1 Adjusting Magnetic Coupling of Overhead Drive

Use 1.5 mm Allen key to loosen set screw on magnetic adapter.

Adjust height to avoid contact with plastic coupling (too low) or coupling interruptions (too high).

Ensure magnetic adapter does not rub against bioreactor coupling.

Tighten set screw after adjustment.

## Notes

Handle bioreactors and sensors carefully due to fragile components and sensitive membranes.

Bioreactor top-heaviness increases with exhaust condenser and overhead drive assembly—ensure placement stability.

Avoid operation outside application limits, especially temperature and pressure.

Avoid contact with concentrated acids, bases, specific antifoams, solvents, and low chemical resistance substances as detailed.

Store bioreactors in original packaging, upright, at 15–25 °C, avoiding UV and strong temperature fluctuations.

Maximum process temperature: 40 °C (BioBLU 0.3c), 45 °C (BioBLU 0.3f).

Maximum contact temperature (heat blanket/wells): 60 °C.

Maximum pressure: 0.4 bar, minimum pressure: 0.0 bar (avoid negative pressure).

Use recommended antifoam agents based on siloxanes.

Perform all sensor calibrations prior to installation.

Work under sterile conditions to maintain bioreactor sterility.

## Review Checklist

[ ] Growth conditions verified (temperature, CO2/O2, time)

[ ] Media composition and volumes verified

[ ] Mixing parameters verified (rpm/RCF, time)

[ ] Incubation times verified

[ ] Centrifugation settings verified (RCF, time, temp) [CHECK]

[ ] Step order confirmed against source

[ ] Any ambiguous values flagged with [CHECK]

## Review Flags

# Review Report for Generated Protocol

## Critical Parameters Checklist (from the protocol)

Working volumes:

Cell culture: approx. 100 mL

Microbiology: 65–70 mL

Agitation speed range:

BioBLU 0.3c: 20 – 500 rpm (source)

BioBLU 0.3f: 20 – 2000 rpm (source) [protocol mentions both ranges generally]

Temperature:

Maximum operating temp: 40 °C (0.3c), 45 °C (0.3f)

Maximum contact temp (e.g., heat blanket): 60 °C

Pressure:

Maximum pressure: 0.4 bar

Minimum pressure: 0.0 bar (avoid negative pressure)

DO sensor details: DASGIP D 4.7 mm, L 162 mm

Sampling valve dead volume: first 1 mL discarded on repeat sampling

Storage conditions: 15 – 25 °C, humidity up to 80% RH (non-condensing)

Magnetic coupling adjustment: set screw with 1.5 mm Allen key

Calibration references: pH sensor calibration per DASware Control User Manual; DO calibration mentioned

Use of ethanol for sterilization/disinfection steps

## Potential Missing Parameters (compare to source)

Exact speed limits per bioreactor version not distinctly specified in the protocol (it states ranges but does not clarify software-limiting differences per version).

The protocol does not explicitly mention working volume minimums (source states 100 - 250 mL for cell culture, 65 - 250 mL for microbiology).

No numeric details on gas flow rates for gassing or headspace overlay flows (source references gas flow management for preventing negative pressure with pump flow rate + 10%, but exact flow values or recommendations missing).

No mention of maximum recommended operation duration (14 days) noted in source specs.

No explicit numeric values for DO sensor insertion depth or membrane stretch (2-3 mm stretch mentioned in source but not detailed in protocol).

No specific instructions on how to handle or limit UV exposure numerically (only general warnings).

No specification of sterilization dose (kGy) or shelf life expiry durations, only generic reference to label.

No explicit mention of pressure limiting devices or overpressure valve settings (only recommendations).

No precise instructions on inoculum volume or stirring ramping profiles.

Centrifugation parameters are not present (mentioned only in checklist as [CHECK]).

## Possible Unsupported Claims

None detected. The protocol closely follows and cites source text sections without adding unsupported information.

## Step Order / Omission Risks

None detected; step orders in the protocol align with source sections 4.2.1 and related subsections.

The protocol includes cautions about sensor membrane damage and emphasizes sterile technique consistent with the source.

The detail about removing dead volume before repeat sampling is consistent and well placed.

Inclusion of warnings about top-heaviness and stability precautions occur in correct order.

No obvious omission of critical assembly or operation steps found.

---

# Summary

The generated protocol comprehensively covers the main parameter values and SOP steps present in the source text with appropriate warnings and cautions. A few numeric details (e.g., exact minimum volumes, gas flow rates, use duration limits) and some measured insertion specifics are not given in numeric detail, which could be important for precise operation in some contexts. These do not pose immediate safety risks if operators consult the full manual but should be noted as areas for further detail.

No unsupported claims or incorrect step sequences were detected. The protocol is suitable for instructional use with the caveat that the flagged numeric and procedural details may require explicit confirmation or supplementation for rigorous process control.